## Slide 1: Use of Sontek RS Live for Data Collection and Review
David S. Mueller
Office of Surface Water
July 2011

## Slide 2: Agenda
Hardware (briefly)
Measurement procedures
Working with the software
Data review
Mobile software (briefly)

## Slide 3: SonTek RiverSurveyor
S5
M9
M9 - 9-beam 3-frequency profiler
  4 beams at 1 Mhz, 4 beams at 3 Mhz
  1 beam @ 500 kHz (depth)
  Velocity profiling range: 1-130 ft (BT)
S5 – 5-beam 2-frequency profiler
  4 beams at 3 Mhz
  1 beam @ 1 MHz (depth)
  Velocity profiling range: 1-17 ft (BT)
Uses RiverSurveyor Live for PC or Mobile for data collection and processing

## Slide 4: Multi-Frequency Multi-Ping
Automatic frequency, ping type (mode), and cell size switching.
Ping types include:
  pulse coherent shallow-water profiling
  pulse incoherent water profiling and bottom-track
  SmartPulseHD – uses pulse-coherent, broadband, and incoherent for precise measurements
  vertical beam depth

## Slide 5: Hydroboard Options
Short (without white line recalled)
Long (all recalled)

## Slide 6: OceanScience
Better in high velocities, but also bigger and heavier
Larger boat above has worked well in > 15 ft/sec

## Slide 7: SonTek Boat Mount Adapter

## Slide 8: Measurement Procedure
Ports / Comm.
Units
GPS Option
Site Information
Date / Time
System Test
Compass Cal.
Magnetic Variation
Transducer Depth
Track Reference
Depth Reference
Salinity
Moving-Bed Test
Start System
Start Edge
Start Moving
End Edge
End Transect
Repeat 15-18

## Slide 9: 1. Ports / Comm.
Direct, Radio or integrated laptop BlueTooth
Supplied BlueTooth connection
Connects at 57600 baud

## Slide 10: Connected to the System

## Slide 11: 2. Units

## Slide 12: 3. GPS Option

## Slide 13: 4. Site Information

## Slide 14: 5. Date / Time

## Slide 15: 6. System Test

## Slide 16: 7. Compass Calibration

## Slide 17: Compass Calibration
Requires two complete circles while pitching and rolling the ADCP and boat significantly.
A calibration with minimal pitch and roll is acceptable if only minimal pitch and roll is experienced during the measurement, such as on a manned boat.
Make sure you are away from any source of magnetic interference (Iron bridge, cell phone, watch, belt buckle, etc)
Notes: RiverSurveyor
The RiverSurveyor compass calibration is a multi-tilt calibration that requires two complete circles while the pitching and rolling the ADCP and the boat significantly.
A calibration with minimal pitch and roll may be acceptable if only minimal pitch and roll is experienced during the measurement, such as on a manned boat. However, even in this situation pitching and rolling the boat as much as possible will improve the calibration.
Click Start and rotate ADCP through two complete circles, slowly, while varying the roll and pitch angles as much as possible.

## Slide 18: Compass Calibration
The calibration score has two parts
  M is magnetic distortion which should be very low (less than 10), again this may not be the case if there is minimal pitch and roll
  Q is a calibration score which should be very high (scale 1-10)
Calibration file is saved in the following location:
  Documents\SonTek Data\YYYY_MM_DD\CompassCal
Notes: Compass Calibration
After completing two complete rotations press the Stop button.
The calibration routines will calculate a calibration score and provide brief comments on the results of the calibration.
Generally the results should say “Pass”, but the Score is often less than excellent if the pitch and roll were minimal.
The calibration score has two parts:
M is the magnetic distortion which should be very low (less than 10)
Q is a calibration score which should be very high on a scale of 1-10

## Slide 19: System Settings

## Slide 20: 13. Moving-Bed Test
Flow
DUP

## Slide 21: Making a Measurement

## Slide 22: Collecting a Transect
Start Edge/
End Edge
Transect
Left Bank
Right Bank
End Edge/
Start Edge
Collect start edge profiles at left/right bank
Start moving and collect transect data
Collect end edge profiles at the opposite bank
Repeat to make additional transects
F5
F5
F5
F5

## Slide 23: Collecting a Transect

## Slide 24: Collecting a Transect
Start Moving/Transect
Select Start Moving(F5) or press F5 and it switches to the Transect tab
Stays on the Transect tab during data collection

## Slide 25: Watch Status Icons
Displays current status information
Blue is generally best
Yellow – caution – (Note Yellow Feet = bad bottom track and will use previous boat speed for 9 samples)
Red – Bad – when feet turn red = bad BT and depth

## Slide 26: Review Track Reference Time Series
3 MHz BT
VTG
1 MHz BT
VTG
1 MHz BT
VTG
VTG
VTG
View if/when track references are being substituted
Area for any samples with track reference “0-Invalid” > 9 times in a row will not have a measured discharge (uses previous good up to 9 times)

## Slide 27: Watch SNR Profiles
If beams separate, lower transducer

## Slide 28: Ending a Measurement
Ending a Measurement:
On completion of the data collection process, click Abort(F8) or press F8
To abort the measurement at any time, click Abort(F8) or press F8

## Slide 29: Where are the data?
System Test – Always on PC or Handheld
Compass Cal. – Always on PC or Handheld
GPS Data – integrated in transect data
Transect Data
  Always on ADCP memory card
  Version 2.5 or later also on PC or Handheld
  Automatically checked for communications errors
  *r.rivr are files logged on PC or Handheld
  *.riv are files logged on ADCP
  Currently references cannot be changed during data collection (Must reprocess transects to change reference)

## Slide 30: Downloading Data
Software warns during data collection of missing samples and download needed
Downloading data is quickest when directly connected to ADCP and using USB-Serial adapter.
Format ADCP routinely after any needed data is downloaded and archived

## Slide 31: RiverSurveyor Live
Discharge
Measurement
Summary
Processing
Toolbox
Sidebar
Tabs to select data file
Tab Window
Display
Tab Window selection
Main Toolbar

## Slide 32: Main Toolbar
“Open File” and “Connect to System”
Each file (transect) or system connection gets its own tab
Swap between tabs by clicking or pressing Ctrl+Tab or Ctrl+Shift+Tab
Tabs may be closed by pressing Ctrl+F4 or clicking the small X in the right corner (in line with tabs – not the Close Program icon)

## Slide 33: Main Toolbar
“Show/Hide Processing Toolbox” and
“Show/Hide Discharge Summary”
Hotkeys
Processing Toolbox (Ctrl+T)
Discharge Summary (Ctrl+S)

## Slide 34: Main Toolbar
Load a loop test (NEW)
Same warning messages as LC

## Slide 35: Tab Windows
Tab Windows show different views of the data
Changing tabs changes all open files
The Sidebar is common to all Tab Windows
Swap tabs by clicking or pressing PageUp and PageDown

## Slide 36: Tab Windows - System
System Tab
A summary all the key settings
Used for real-time setup
Can be printed
Items in green have been changed since the field data collection.

## Slide 37: Tab Windows - Samples
Samples Tab
Tabular display for all samples
Contour and additional profile graph
Tabular can be customized
Use arrows to scroll up or down and left and right arrows to jump to top (beginning) or bottom (end) of data
Edges are displayed in the tabular data

## Slide 38: Tab Windows - Navigation
Navigation Tab
Good all purpose
Track and Contour
Edges are not displayed here

## Slide 39: Tab Windows - Edges
Edges Tab
Track, Profile and
Contour for each edge
Good tool to evaluate:
Variation in positions
Variation in depth
Relative position between references
Only Edges are displayed in this tab

## Slide 40: Tab Windows – Time Series
Time Series Tab
Multiple time series graphs at once
Edges are not displayed here

## Slide 41: Tab Windows - Transect
Transect Tab
Good for comparisons of parameters over the length of the transect
Time Series, Track, Contour
Edges are not displayed here

## Slide 42: Graphs
Right click on graph to change what is displayed.

## Slide 43: Graphs – Track
Show Velocity Vectors
Hides vectors to make it
easier to see the tracks
Controlling the Track
Left click and drag to move the graph around
Use mouse scroll wheel to zoom in or out or use Z and Shift+Z keys
Change vector length using Ctrl++ or Ctrl+- (AutoScale must be off)

## Slide 44: Processing Toolbox
All processing controls are here
Only available during post-processing
Some parameters (System and Measurement settings) can be modified on the System tab
Show/Hide Processing Toolbox using icon in main toolbar
Tool Sets selected using combo box
Updates with each file selected
Lock transect
Reset to Field icon

## Slide 45: Processing Toolbox
Matlab Export
Used to create files that are for importing into USGS programs (LC, SMBA, extrap, EDI)
Contains all data and settings
Parameters are grouped
Velocities are in the reference and coordinate system at the time of export

## Slide 46: Processing Toolbox – System Settings
Adjust field settings as necessary
Screening distance – distance below transducer for valid data. (Normally set to zero, use default blanking distance)
Combinations of references and coordinates must be logical i.e. GPS in beam coordinates does not make a lot of sense

## Slide 47: Edge Estimates
Sloped Bank
Vertical Bank
User Input (estimated Q)
  Must document method used to estimate Q
    AA measured depths and velocities
    Visual estimate of depths and vel.
    Estimated Q between software calculated Q from sloped and vertical
    Other...
Use “Apply to all” for multiple transects from the same measurement

## Slide 48: Setting Extrapolations
Power / Power
  Default
  Change exponent
    Must change for both top and bottom
Constant / Power
  Set top “Method” to constant
  Set top “Cells to use” to “Only Use”
  Set top “cells” to 1
Constant / No Slip
  Set top “Method” to constant
  Set top “Cells to use” to “Only Use”
  Set top “cells” to 1
  Set bottom “Cells to use” to “Only Use”
  Set bottom “cells” to 2 or 1 (generally recommend 2)

## Slide 49: Processing Toolbox – Thresholds
Heading correction – allows manual compensation for poor compass calibration
Minimum SNR – used for data screening (recommend not changing)
Cell Screening – turns of screening of data below sidelobe cutoff.
Composite Tracks – turns on and off automatic substitution of GPS and BT.

## Slide 50: Discharge Measurement Summary
Toolbar Buttons
Summary Report
ASCII Export
Delete Selected (highlighted) Records
Delete All Records
Highlight outside 5%

## Slide 51: Discharge Measurement Summary
Full color report
Based on past report structure
Site Information is drawn from data files
Settings that are not common are highlighted with an asterisk
Unchecked records are not displayed

## Slide 52: Discharge Summary NEW
Prompt for compass calibration
Prompt for system test file
Compass calibration, system test, and loop test included in output

## Slide 53: Review Track Reference
Determine best track reference
Look at ship track plot for any irregularities
  Uncheck composite tracks to verify reference not substituted
  Typically check composite tracks for final Q

## Slide 54: Significant Bad Bottom Track
Ref: BT
Composite Track Off
No Q or width Computed when Track Ref Invalid (Boat speed = 0)
Ref: BT
Composite Track On
VTG used when BT invalid
Shiptrack has VTG data plotted with Bottom-Track (Blue)
This measurement had moving bed and good GPS data, so VTG or GGA should be used for track reference

## Slide 55: Determine Best Depth Reference
Look at VB depth and BT Beam Depth on a time series plot… make sure they are plotted at the same scale
Will substitute with alternate reference if selected reference not available (not an option)

## Slide 56: RS Live Mobile
Almost everything in PC software is also in mobile Software
Must have Bluetooth PCM
Loop Method
NEW
Some reported issues of locking up after loop

## Slide 57: Use of SonTek S5/M9
Make comparison measurements for the range of conditions you intend to use the instrument
  Water velocities
  Streambed type
  Flow depths
  Turbulence
  Sediment concentration
You can then use the instrument if there are no significant deviations from the comparison measurements
  Consistent biases are of particular concern
If a check measurement is indicated, use a different instrument
Document and submit comparisons to OSW to assist in finalizing the evaluation of these instruments. https://collaboration.usgs.gov/wg/oswha/Testing/default.aspx

## Slide 58: Questions?